TCG NEBS COMPLIANCE TEST Report FOR:

<Name of Company>

<Company Address>

Product: <Product Name>

Part 10

Electromagnetic interference (EMI)

Section 3, GR-1089-CORE

Telcordia Technologies GR-1089, Issue 6, May 2011

Date: <<date>>

Project: <Project #, if applicable>

Job: <Job #, if applicable>

Report: <Report #>

TABLE OF CONTENTS

Test Results Summary	3

Overview	5

FCC Criteria (3.1.3)	6

Documentation (3.1.5)	7

Electric Fields Radiated Emissions Criteria For Unitentional Radiators (3.2.1.1)	8

Electric Fields Radiated Emission Criteria for Intentional Radiators (3.2.1.2)	11

Conducted Emissions Requirements On AC Power Ports – Voltage (3.2.2.1)	13

Conducted Emission Requirements for AC and DC Power - Current (3.2.2.2)	15

Conducted Emission Requirements for Telecommunications Ports– Current (3.2.3)	17

Electric Fields Radiated Immunity Criteria with Doors and Covers Closed (3.3.1.1)	18

Electric Fields Radiated Immunity Criteria with Doors and Covers Open (3.3.1.2)	20

Conducted Immunity Criteria for AC and DC Power Ports (3.3.2)	22

Conducted Immunity Requirements for Signal and telecommunications Leads (3.3.3.1)	24

Conducted Immunity Requirements for Voiceband Leads (3.3.3.2)	26

Appendix A	28

- LIST OF TABLES

Page

Table 10-1: Electromagnetic Interference Summary of Test Results	3

Table 10-2: Radiated E-Field Detailed List of Test Equipment	10

Table 10-3: Radiated E-Field (Intentional Radiators) Detailed List of Test Equipment	12

Table 10-4: Conducted Voltage Emissions on AC Power Detailed List of Test Equipment	14

Table 10-5: Conducted Current Emissions on AC/DC Power Detailed List of Test Equipment	16

Table 10-9: Electric Field Radiated Immunity with Doors Closed Detailed List of Test Equipment	19

Table 10-10: Electric Field Radiated Immunity with Doors and Covers Open Detailed List of Test Equipment	21

Table 10-11: Conducted Immunity for AC and DC Power Ports Detailed List of Test Equipment	23

Table 10-12: Conducted Immunity for Signal Detailed List of Test Equipment	25

Table 10-13: Conducted Immunity for Voiceband and Telecommunications Leads Detailed List of Test Equipment	27

LIST OF FIGURES

Page

Error! No table of figures entries found.

## Test Results Summary

<<Summarize the test results here, including hardware changes or changes in product configuration during the testing.  While this is a summary, it should contain most of the relevant details to the tests and the test results. >>

### Column Heading Definitions for Summary of Test Results Table

The following Summary of Test Results table contains these columns of information:

- Section column gives the Section numbers from GR-1089-CORE.

- Section Name column gives the Section name from GR-1089-CORE.

- Criteria column gives the local number of the requirement (e.g., R3-1) from GR-1089-CORE and the absolute number of the requirement (e.g., [2]).

- Results column gives the results of the evaluation (Compliant, Non-compliant, etc.).

- Compliant: The Equipment Under Test met the requirements of the corresponding criteria.

- Non-compliant: The Equipment Under Test did not meet the requirements of the corresponding criteria.

- NA: The criteria were Not Applicable to Equipment Under Test {Explanation Required}

- ENR: An Evaluation, to these criteria, was Not Requested by the customer.

- Comments column gives the page num

- Page column gives the page number, in this report, for the corresponding criteria.

Table 10-1: Electromagnetic Interference Summary of Test Results

| ##### Section | ##### Section Name | ##### Criteria | ##### Results #### Pass/ Fail/ NA | ##### Comments | ##### Page |
| --- | --- | --- | --- | --- | --- |
| 3. | Electromagnetic Interference | - | - |  | - |
| 3.1.5 | Documentation | R3-1 [159] |  |  | 7 |
| 3.2.1.1 | Electric Fields Radiated Emission Criteria for Unintentional Radiators | R3-2 [8] |  |  | 8 |
|  |  | O3-3 [9] |  |  | 8 |
|  |  | R3-4 [10] |  |  | 8 |
| 3.2.1.2 | Electric Fields Radiated Emission Criteria for Intentional Radiators | R3-5 [160] |  |  | 11 |
|  |  | R3-6 [161] |  |  | 11 |
|  |  | R3-7 [162] |  |  | 11 |
|  |  | O3-8 [163] |  |  | 11 |
| 3.2.2.1 | Conducted Emission Requirements for AC Power Ports – Voltage | R3-9 [12] |  |  | 13 |
| 3.2.2.2 | Conducted Emission Requirements for AC and DC Power Ports – Current | R3-10 [13] |  |  | 15 |
| 3.2.3 | Conducted Emission Criteria for Telecommunications Ports - Current | R3-11 [128] |  |  | 17 |
| 3.3.1 | Radiated Immunity Criteria | - |  |  |  |
| 3.3.1.1 | Electric Fields Radiated Immunity Criteria with Doors and Covers Closed | R3-12 [16] |  |  | 18 |
| 3.3.1.2 | Electric Fields Radiated Immunity Criteria with Doors and Covers Open | R3-13 [18] |  |  | 20 |
| 3.3.2 | Conducted Immunity Criteria for AC and DC Power Ports | R3-14 [130] |  |  | 22 |
| 3.3.3.1 | Conducted Immunity Requirement for Signal and Telecommunications Leads | R3-15 [131] |  |  | 24 |
| 3.3.3.2 | Conducted Immunity Requirements for Voiceband Leads | R3-16 [164] |  |  | 26 |

NA: Not Applicable; ENR: Evaluation Not Requested

## Overview

### Project Objective

Testing was performed to determine if the <Product Name> met the requirements for Section 3, Electromagnetic Interference, of Telcordia Technologies GR-1089-CORE, Issue 6, May 2011.

### Equipment Configuration and Test Conditions

<Specify number of samples tested>

<<Describe EUT Configuration using photographs, diagrams, words, or any combination thereof>>

<<Detailed pack-fill information should be entered in a table in Appendix A at the end of this section>>

### Equipment Operating Conditions

<<Place pictures and diagrams of EUT operational set-up>>

### Pass/ Fail Criteria

<<Provide pass/fail criteria for the test>>

## FCC Criteria (3.1.3)

“The FCC defines a Class-A digital device as one marketed for use in a commercial, industrial or business environment, exclusive of a device which is marketed for use by the general public or is intended to be used in the home. Most equipment located in dedicated telephone company premises should be considered to be Class-A digital devices and shall meet all requirements specified by the FCC for such devices. Over the FCC frequency range, the emission limits in this GR are the same as those specified by the FCC. The FCC defines a Class-B digital device as one marketed for use in a residential environment notwithstanding use in commercial, business and industrial environment. These include personal computers, calculators, and similar electronic devices for use by the general public.”

<<Describe the intended use and environment, and hence the applicable FCC Class of the EUT>>

## Documentation (3.1.5)

### Criteria:

R3-1 [159] 	The installation or instruction manual shall specify any special accessories such as shielded cables or special connectors that are used to achieve compliance for emission and immunity criteria.

NOTE: See FCC Part 15.27, [12] Special Accessories, for more details.

### Test Location:

The following evaluation was performed by (Engineer Name) between/on date(s) at the < ITL Name and Address>.

### Test Method:

<<Describe test method in detail>>

<< Include any additional test conditions (e.g. signal and power cable dressing) from the NEBS Compliance Checklist which pertain to the electromagnetic compatibility and electrical safety criteria>>

<<State if testing to the GR-1089-CORE test method and describe any deviations from the GR documented test method>>

### Test Results:

The <Product Name> was compliant/ not compliant with Requirement R3-1 [159].

<<Describe test results in detail where relevant. Include graphs, tables, scans, chart recorder data, etc. of test results>>

## Electric Fields Radiated Emissions Criteria For Unitentional Radiators (3.2.1.1)

### Criteria:

R3-2 [8]	Radiated emissions from the EUT shall not exceed the levels of field strength given in Table 3-1. An average detector shall be used for frequencies above 1 GHz. A quasi-peak detector shall be used for frequencies in the range of 30 MHz to 1 GHz. This test shall be done with all EUT doors and covers closed. For equipment that does not have doors or covers R3-2 [8] shall be applied.

O3-3 [9]	Radiated emissions from the EUT should not exceed the levels of field strength given in Table 3-1 with all doors and covers open. Doors and covers that are not intended to be opened during operation, maintenance, and/or repair need not be opened.

R3-4 [10]	Radiated emissions from Class A and B EUTs shall not exceed the levels of field strength given in Table3-2. An average detector shall be used for frequencies below 150 kHz and for frequencies above 1 GHz. A quasi-peak detector shall be used for frequencies in the range of 150 kHz to 1 GHz. This test shall be done with all the EUT doors and covers open. Doors and covers that are not intended to be opened during EUT operation, maintenance, and/or repair need not be opened.

(Tables from GR-1089)

### Test Location:

The following evaluation was performed by (Engineer Name) between/on date(s) at the < ITL Name and Address>.

### Test Method:

<<Describe test method in detail>>

<< Include any additional test conditions (e.g. signal and power cable dressing) from the NEBS Compliance Checklist which pertain to the electromagnetic compatibility and electrical safety criteria>>

<<State if testing to the GR-1089-CORE test method and describe any deviations from the GR documented test method>>

### Test Results:

The <Product Name> is compliant/ not compliant with Requirement R3-2 [8].  No emissions originating from the equipment were above the limit shown in Table 3-1.  By demonstrating compliance/ non-compliance with Objective O3-3 [9], the <Product Name> is considered compliant/ not compliant with Requirements R3-2 [8] and R3-4 [10].

<<Describe test results in detail where relevant. Include graphs, tables, scans, chart recorder data, etc. of test results>>

### Test Equipment Used

Table 10-2: Radiated E-Field Detailed List of Test Equipment

| Description | Serial Number | Manufacturer | Model | Cal. Date | Cal. Due |
| --- | --- | --- | --- | --- | --- |

<<Fill out the table of test equipment used for this test ONLY>>

## Electric Fields Radiated Emission Criteria for Intentional Radiators (3.2.1.2)

### Criteria:

R3-5 [160] 	The EUT shall be certified to applicable FCC parts in 47 CFR.

R3-6 [161] 	Radiated emissions from the EUT shall not exceed the levels of field strength provided in Table 3-1. The limits apply to the frequencies of unwanted emissions and not the fundamental frequency. An average detector shall be used for frequencies above 1 GHz. A quasi-peak detector shall be used for frequencies in the range of 30 MHz to 1 GHz. This test shall be done with all EUT doors and covers closed.

R3-7 [162] 	When testing the electric-field radiated emissions of the EUT, the emissions at the fundamental frequency of the EUT shall be measured and provided in a test report.

O3-8 [163] 	Radiated emissions from the EUT should not exceed the levels of field strength given in Table3-1 for Class A with all doors and covers open. The limits apply to the frequencies of unwanted emissions and not the fundamental frequency. Doors and covers that are not intended to be opened during operation, maintenance, and/or repair need not be opened.

(Table from GR-1089)

### Test Location:

The following evaluation was performed by (Engineer Name) between/on date(s) at the <ITL Name and Address>.

### Test Method:

<<Describe test method in detail>>

<< Include any additional test conditions (e.g. signal and power cable dressing) from the NEBS Compliance Checklist which pertain to the electromagnetic compatibility and electrical safety criteria>>

<<State if testing to the GR-1089-CORE test method and describe any deviations from the GR documented test method>>

### Test Results:

The <Product Name> was compliant/ not compliant with Requirement R3-5 [160] and met the applicable FCC parts in 47 CFR. The <Product Name> is compliant/ not compliant with Requirement R3-6 [161].  The <Product Name> was/was not tested to at the fundamental frequency of the EUT and the data was/was not reported which demonstrated compliance/ non- compliance with Requirement R3-7 [162]. The <Product Name> was compliant/ not compliant with Requirement R3-8 [163]. No emissions originating from the equipment were above the limit shown in Table 3-1 (Class A) with all the doors and covers opened.

<<Describe test results in detail where relevant. Include graphs, tables, scans, chart recorder data, etc. of test results>>

### Test Equipment Used

Table 10-3: Radiated E-Field (Intentional Radiators) Detailed List of Test Equipment

| Description | Serial Number | Manufacturer | Model | Cal. Date | Cal. Due |
| --- | --- | --- | --- | --- | --- |

<<Fill out the table of test equipment used for this test ONLY>>

<<Fill out the table of test equipment used for this test ONLY>>

## Conducted Emissions Requirements On AC Power Ports – Voltage (3.2.2.1)

### Criteria:

R3-9 [12]	Conducted emissions from the EUT into low-voltage ac public utility power lines shall not exceed the average voltage limits and quasi-peak voltage limits given in Table 3-3 for Class A EUT or Table 3-4 for Class B EUT when using an average and a quasi-peak detector, respectively, and measured in accordance with Section3.4.5.1.1.

In Table 3-3 and Table 3-4, the frequency is that of the emission under investigation, and the stricter limit applies at the edge between the two frequency bands.

(Tables from GR-1089)

### Test Location:

The following evaluation was performed by (Engineer Name) between/on date(s) at the <ITL Name and Address>.

### Test Method:

<<Describe test method in detail>>

<< Include any additional test conditions (e.g. signal and power cable dressing) from the NEBS Compliance Checklist which pertain to the electromagnetic compatibility and electrical safety criteria>>

<<State if testing to the GR-1089-CORE test method and describe any deviations from the GR documented test method>>

### Test Results:

The <Product Name> was compliant/ not compliant with Requirement R3-9 [12].

<<Describe test results in detail where relevant. Include graphs, tables, scans, chart recorder data, etc. of test results>>

### Test Equipment Used

Table 10-4: Conducted Voltage Emissions on AC Power Detailed List of Test Equipment

| Description | Serial Number | Manufacturer | Model | Cal. Date | Cal. Due |
| --- | --- | --- | --- | --- | --- |

<<Fill out the table of test equipment used for this test ONLY>>

## Conducted Emission Requirements for AC and DC Power - Current (3.2.2.2)

### Criteria:

R3-10 [13]	Conducted common-mode emissions from the EUT into low-voltage ac public utility power lines and dc power conductors shall not exceed the levels given in Table 3-5 measured in accordance with Section 3.4.5.1.2 and Section 3.4.5.1.3. An average detector shall be used for frequencies below 150 kHz. A quasi-peak detector shall be used for frequencies above 150 kHz.

(Tables from GR-1089)

### Test Location:

The following evaluation was performed by (Engineer Name) between/on date(s) at the < ITL Name and Address>.

### Test Method:

<<Describe test method in detail>>

<< Include any additional test conditions (e.g. signal and power cable dressing) from the NEBS Compliance Checklist which pertain to the electromagnetic compatibility and electrical safety criteria>>

<<State if testing to the GR-1089-CORE test method and describe any deviations from the GR documented test method>>

### Test Results:

The <Product Name> was compliant/ not compliant with Requirement R3-10 [13].

<<Describe test results in detail where relevant. Include graphs, tables, scans, chart recorder data, etc. of test results>>

### Test Equipment Used

Table 10-5: Conducted Current Emissions on AC/DC Power Detailed List of Test Equipment

| Description | Serial Number | Manufacturer | Model | Cal. Date | Cal. Due |
| --- | --- | --- | --- | --- | --- |

<<Fill out the table of test equipment used for this test ONLY>>

## Conducted Emission Requirements for Telecommunications Ports– Current (3.2.3)

Criteria:

R3-11 [128]    Conducted common-mode emissions from the EUT into signal, telecommunication, and analog voiceband leads shall not exceed the levels given in Table 3-6 measured in accordance with Section 3.4.5.1.4. An average detector shall be used for frequencies below 150 kHz. A quasi-peak detector shall be used for frequencies above 150 kHz.

(Table from GR-1089)

### Test Location:

The following evaluation was performed by (Engineer Name) between/on date(s) at the < ITL Name and Address>.

### Test Method:

<<Describe test method in detail>>

<< Include any additional test conditions (e.g. signal and power cable dressing) from the NEBS Compliance Checklist which pertain to the electromagnetic compatibility and electrical safety criteria>>

<<State if testing to the GR-1089-CORE test method and describe any deviations from the GR documented test method>>

### Test Results:

The <Product Name> was compliant/ not compliant with Requirement R3-12 [16].

<<Describe test results in detail where relevant. Include graphs, tables, scans, chart recorder data, etc. of test results>>

## Electric Fields Radiated Immunity Criteria with Doors and Covers Closed (3.3.1.1)

### Criteria:

R3-12 [16]	The EUT shall perform properly when subjected to the calibration electric fields of 138.6 dBμV/m (8.5 V/m), RMS value of the unmodulated test signal over the frequency range of 0.01 through 10,000 MHz. During the test, the test signal shall be modulated in accordance with Section 3.5.5.3.6. In general, performance criteria are given in the applicable Telcordia Generic Requirements, national and international standards; examples of particular performance criteria are found in Section3.5.3. This test is to be performed with all the EUT doors and covers closed.

### Test Location:

The following evaluation was performed by (Engineer Name) between/on date(s) at the < ITL Name and Address>.

### Test Method:

<<Describe test method in detail>>

<< Include any additional test conditions (e.g. signal and power cable dressing) from the NEBS Compliance Checklist which pertain to the electromagnetic compatibility and electrical safety criteria>>

<<State if testing to the GR-1089-CORE test method and describe any deviations from the GR documented test method>>

### Test Results:

The <Product Name> was compliant/ not compliant with Requirement R3-12 [16].

<<Describe test results in detail where relevant. Include graphs, tables, scans, chart recorder data, etc. of test results>>

### Test Equipment Used

Table 10-6: Electric Field Radiated Immunity with Doors Closed Detailed List of Test Equipment

| Description | Serial Number | Manufacturer | Model | Cal. Date | Cal. Due |
| --- | --- | --- | --- | --- | --- |

<<Fill out the table of test equipment used for this test ONLY>>

## Electric Fields Radiated Immunity Criteria with Doors and Covers Open (3.3.1.2)

### Criteria:

R3-13 [18]	The EUT shall perform properly when it is subjected to calibration electric field of 138.6 dBμV/m (8.5 V/m), RMS value of the unmodulated test signal over the frequency range of 0.01 through 10,000 MHz. During the test, the test signal shall be modulated in accordance with Section 3.5.5.3.6. In general, performance criteria are found in the applicable Telcordia Generic Requirements, national and international standards; examples of particular performance criteria are found in Section3.5.3. This test shall be done with all the EUT doors and covers open, except for doors and covers that are not intended to be opened during EUT operation, maintenance, and/ or repair, which may remain closed.

### Test Location:

The following evaluation was performed by (Engineer Name) between/on date(s) at the < ITL Name and Address>.

### Test Method:

<<Describe test method in detail>>

<< Include any additional test conditions (e.g. signal and power cable dressing) from the NEBS Compliance Checklist which pertain to the electromagnetic compatibility and electrical safety criteria>>

<<State if testing to the GR-1089-CORE test method and describe any deviations from the GR documented test method>>

### Test Results:

The <Product Name> was compliant/ not compliant with Requirement R3-13 [18].

<<Describe test results in detail where relevant. Include graphs, tables, scans, chart recorder data, etc. of test results>>

### Test Equipment Used

Table 10-7: Electric Field Radiated Immunity with Doors and Covers Open Detailed List of Test Equipment

| Description | Serial Number | Manufacturer | Model | Cal. Date | Cal. Due |
| --- | --- | --- | --- | --- | --- |

<<Fill out the table of test equipment used for this test ONLY>>

## Conducted Immunity Criteria for AC and DC Power Ports (3.3.2)

### Criteria:

R3-14 [130]	The EUT shall perform properly when electromagnetic energy is injected longitudinally (common-mode) into each power port of the EUT, one port at a time, at the previously calibrated limits given in Table 3-7. If the EUT fails the pre- calibrated limit of Table 3-7 because of excessive coupling, the injected current into the EUT leads may be adjusted to a level 6 dB above the calibration current of Table 3-7. If the EUT performs properly after the current is adjusted, then the EUT is considered to have met the criteria. In general, performance criteria are found in the applicable Telcordia Generic Requirements, national and international standards. Examples of particular performance criteria are found in Section 3.5.3. The lower limit applies at the band edges. In Table 3-7, f is the frequency of the injected electromagnetic energy in MHz.

(Table from GR-1089)

### Test Location:

The following evaluation was performed by (Engineer Name) between/on date(s) at the < ITL Name and Address>.

### Test Method:

<<Describe test method in detail>>

<< Include any additional test conditions (e.g. signal and power cable dressing) from the NEBS Compliance Checklist which pertain to the electromagnetic compatibility and electrical safety criteria>>

<<State if testing to the GR-1089-CORE test method and describe any deviations from the GR documented test method>>

### Test Results:

The <Product Name> was compliant/ not compliant with Requirement R3-14 [130].

<<Describe test results in detail where relevant. Include graphs, tables, scans, chart recorder data, etc. of test results>>

### Test Equipment Used

Table 10-8: Conducted Immunity for AC and DC Power Ports Detailed List of Test Equipment

| Description | Serial Number | Manufacturer | Model | Cal. Date | Cal. Due |
| --- | --- | --- | --- | --- | --- |

<<Fill out the table of test equipment used for this test ONLY>>

## Conducted Immunity Requirements for Signal and telecommunications Leads (3.3.3.1)

### Criteria:

R3-15 [131]	The EUT shall perform properly when electromagnetic energy is injected longitudinally (common-mode) into signal leads associated with each kind of port of the EUT, one kind of port at a time, at the previously calibrated limits given in Table 3-7. At least one of the identical ports in the EUT shall be tested. If the EUT fails the pre-calibrated limit of Table 3-7 because of excessive coupling, the injected current into the EUT leads may be adjusted to a level 6 dB above the calibration current of Table 3-7. If the EUT performs properly after the current is adjusted, then the EUT is considered to have met the criteria. In general, performance criteria are found in the applicable Telcordia Generic Requirements, national and international standards; examples of particular performance criteria are found in Section3.5.3. In Table 3-7, f is the frequency of the injected electromagnetic energy in MHz.

### (Table from GR-1089)

### Test Location:

The following evaluation was performed by (Engineer Name) between/on date(s) at the < ITL Name and Address>.

### Test Method:

<<Describe test method in detail>>

<< Include any additional test conditions (e.g. signal and power cable dressing) from the NEBS Compliance Checklist which pertain to the electromagnetic compatibility and electrical safety criteria>>

<<State if testing to the GR-1089-CORE test method and describe any deviations from the GR documented test method>>

### Test Results:

The <Product Name> was compliant/ not compliant with Requirement R3-15 [131].

<<Describe test results in detail where relevant. Include graphs, tables, scans, chart recorder data, etc. of test results>>

### Test Equipment Used

Table 10-9: Conducted Immunity for Signal Detailed List of Test Equipment

| Description | Serial Number | Manufacturer | Model | Cal. Date | Cal. Due |
| --- | --- | --- | --- | --- | --- |

<<Fill out the table of test equipment used for this test ONLY>>

## Conducted Immunity Requirements for Voiceband Leads (3.3.3.2)

### Criteria:

R3-16 [164]	The EUT shall perform properly when electromagnetic energy is injected longitudinally (common-mode) into each port of associated voiceband and telecommunications leads of the EUT, one port at a time, at the previously calibrated limits given in Table 3-8. If the EUT fails the pre-calibrated limit of Table 3-8 because of excessive coupling, the injected current into the EUT leads may be adjusted to a level 6 dB above the calibration current of Table3-8. If the EUT performs properly after the current is adjusted, then the EUT is considered to have met the criteria. In general, performance criteria are found in the applicable Telcordia Generic Requirements, national and international standards; examples of particular performance criteria are found in Section 3.5.3. In Table 3-8, f is the frequency of the injected electromagnetic energy in MHz.

(Table from GR-1089)

### Test Location:

The following evaluation was performed by (Engineer Name) between/on date(s) at the < ITL Name and Address>.

### Test Method:

<<Describe test method in detail>>

<< Include any additional test conditions (e.g. signal and power cable dressing) from the NEBS Compliance Checklist which pertain to the electromagnetic compatibility and electrical safety criteria>>

<<State if testing to the GR-1089-CORE test method and describe any deviations from the GR documented test method>>

### Test Results:

The <Product Name> was compliant/ not compliant with Requirement R3-16 [164].

<<Describe test results in detail where relevant. Include graphs, tables, scans, chart recorder data, etc. of test results>>

### Test Equipment Used

Table 10-10: Conducted Immunity for Voiceband and Telecommunications Leads Detailed List of Test Equipment

| Description | Serial Number | Manufacturer | Model | Cal. Date | Cal. Due |
| --- | --- | --- | --- | --- | --- |

<<Fill out the table of test equipment used for this test ONLY>>

<<Insert tables with detailed description of equipment pack fill>>

Section 3.1.3, FCC Criteria, of GR-1089-CORE, Issue 2, December 1997.